# National Trade Measurement Amendment Regulation 2012 (No. 1)1

Select Legislative Instrument 2012 No. 302

I, QUENTIN BRYCE, Governor-General of the Commonwealth of Australia, acting with the advice of the Federal Executive Council, make the following regulation under the National Measurement Act 1960.

Dated 6 December 2012

QUENTIN BRYCE

Governor-General

By Her Excellency’s Command

GREG COMBET

Minister for Industry and Innovation

1	Name of regulation

This regulation is the National Trade Measurement Amendment Regulation 2012 (No. 1).

2	Commencement

This regulation commences on 1 January 2013.

3	Amendment of National Trade Measurement Regulations 2009

Schedule 1 amends the National Trade Measurement Regulations 2009.

Schedule 1	Amendments

(section 3)

[1]	Paragraph 5.6 (b)

substitute

(b)	electricity meters installed before 1 January 2013;

(ba)	electricity meters installed on or after 1 January 2013, other than electricity meters that measure less than 750 MWh of energy per year;

[2]	Paragraph 5.6 (d)

substitute

(d)	water meters installed on or after 1 July 2004, other than cold water meters with a maximum continuous flow rate capacity of not more than 4 000 litres per hour.

[3]	Schedule 1, Part 3, Division 2, clause 1

substitute

1	For in-service inspection of instruments with digital indication, add 0.5 scale interval to the maximum permissible error for in-service inspection that applies to an analog instrument.

1A	However, item 1 does not apply to an instrument with digital indication if the scale interval for the instrument is less than or equal to 0.2 dm2.

[4]	Schedule 1, Part 3, Division 6, clause 5

substitute

5	The maximum permissible error for any load equal to or greater than the minimum capacity and equal to or less than the maximum capacity in automatic operation is:

(a)	if the national instrument test procedures that apply to catch weighers eliminate the need for digital rounding—the maximum permissible error set out in table 13 minus a verification scale interval of 0.5 e; or

(b)	in any other case—set out in table 13.

Note   The national instrument test procedures are defined in the Act and are available at www.nmi.gov.au.

[5]	Schedule 1, Part 3, Division 11

omit

Table 1	Maximum permissible errors for water meters

insert

Table 17	Maximum permissible errors for water meters

[6]	Schedule 1, Part 3, after Division 11

insert

Division 12	Electricity meters

1	In this Division:

Ib, for an electricity meter of a kind mentioned in clause 3, is the basic current.

Imax, for an electricity meter of a kind mentioned in clause 3 or 4, is the maximum current.

In, for an electricity meter of a kind mentioned in clause 4, is the nominal current.

Accuracy classes

2	Electricity meters are classified into the following accuracy classes:

(a)	0.2;

(b)	0.5;

(c)	1;

(d)	1.5.

Maximum permissible errors—direct-connected electricity meters

3	The maximum permissible errors for the following kinds of electricity meters are set out, for an item, in columns 4 and 5 of table 18:

(a)	single phase direct-connected electricity meters with an accuracy class mentioned, for the item, in column 4 or 5 of that table;

(b)	polyphase direct-connected electricity meters with an accuracy class mentioned, for the item, in column 4 or 5 of that table.

4	The maximum permissible errors mentioned in column 4 or 5 of table 18 apply at the current rate and power factor mentioned, for an item, in column 2 and 3 of that table.

Table 18	Single phase and polyphase direct-connected electricity meters

| Item | Current range | Power factor | Maximum permissible error (%) |  |
| --- | --- | --- | --- | --- |
|  |  |  | Accuracy class 1 | Accuracy class 1.5 |
| 1 | 0.05 Ib ≤ I < 0.1 Ib | 1 | ±1.5% | ±1.5% |
| 2 | 0.1 Ib ≤ I ≤ Imax | 1 | ±1.0% | ±1.5% |
| 3 | 0.1 Ib ≤ I < 0.2 Ib | 0.5 inductive | ±1.5% | ±1.5% |
| 4 | 0.1 Ib ≤ I < 0.2 Ib | 0.8 capacitive | ±1.5% | - |
| 5 | 0.2 Ib ≤ I ≤ Imax | 0.5 inductive | ±1.0% | ±1.5% |
| 6 | 0.2 Ib ≤ I ≤ Imax | 0.8 capacitive | ±1.0% | - |

Maximum permissible errors—transformer-operated electricity meters

5	The maximum permissible errors for the following kinds of electricity meters are set out, for an item, in columns 4, 5 and 6 of table 19:

(a)	single phase transformer-operated electricity meters with an accuracy class mentioned, for the item, in column 4, 5 or 6 of that table;

(b)	polyphase transformer-operated electricity meters with an accuracy class mentioned, for the item, in column 4, 5 or 6 of that table.

6	The maximum permissible errors mentioned in column 4, 5 or 6 of table 19 apply at the current rate and power factor mentioned, for an item, in column 2 and 3 of that table.

Table 19	Single phase and polyphase transformer-operated electricity meters

| Item | Current range | Power factor | Maximum permissible error (%) |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Accuracy class 0.2 | Accuracy class 0.5 | Accuracy class 1 |
| 1 | 0.01 In ≤ I < 0.05 In | 1 | ±0.4% | ±1.0% | - |
| 2 | 0.02 In ≤ I < 0.05 In | 1 | - | - | ±1.5% |
| 3 | 0.05 In ≤ I ≤ Imax | 1 | ±0.2% | ±0.5% | ±1.0% |
| 4 | 0.02 In ≤ I < 0.1 In | 0.5 inductive | ±0.5% | ±1.0% | - |
| 5 | 0.02 In ≤ I < 0.1 In | 0.8 capacitive | ±0.5% | ±1.0% | - |
| 6 | 0.05 In ≤ I < 0.1 In | 0.5 inductive | - | - | ±1.5% |
| 7 | 0.05 In ≤ I < 0.1 In | 0.8 capacitive | - | - | ±1.5% |
| 8 | 0.1 In ≤ I ≤ Imax | 0.5 inductive | ±0.3% | ±0.6% | ±1.0% |
| 9 | 0.1 In ≤ I ≤ Imax | 0.8 capacitive | ±0.3% | ±0.6% | ±1.0% |

[7]	Schedule 2, Part 1, subitems 5.1 to 5.3

substitute

|  | 5.1 | Fuel dispensers used for petroleum products other than LPG |
| --- | --- | --- |
|  | 5.2 | Flow meters used for petroleum products |
|  | 5.3 | Flow meters used for liquids other than petroleum products |

[8]	Schedule 2, Part 1, subitems 10.1 and 10.2

substitute

|  | 10.1 | Fuel dispensers used for LPG, other than cryogenic liquids |
| --- | --- | --- |
|  | 10.2 | Flow meters used for LPG, other than cryogenic liquids |

[9]	Schedule 2, Part 1, subitems 15.1 to 15.3

substitute

|  | 15.1 | Grain |
| --- | --- | --- |
|  | 15.2 | Cane sugar |
|  | 15.3 | Wine grapes |

Note

1.	All legislative instruments and compilations are registered on the Federal Register of Legislative Instruments kept under the Legislative Instruments Act 2003. See www.comlaw.gov.au.